1/30/01

Microscopy

The range of size in parasites is large, varying from 1 micrometer in Leishmania to 37 feet for some tapeworms.  Because nearly all parasites have life cycle stages that are microscopic (many large parasite produce small immature stages), the proper use of the microscope is crucial to successful parasitology.

All microscopes provide two types of power—magnification and resolution (also called resolving power).  Magnification is the ratio between the size of the image and the object.  The total magnifying power is determined by multiplying the power of the objective lens by the power of the ocular lens.  Resolution is a measure of the clarity of the image; it is the minimum distance two points can be separated and still be distinguished as two points.  The function of the microscope is to enable us to discern two points as separate and distinct even when they are so close together that they appear as one to the unaided eye.  This distance at which two points appear as one is about 0.1 mm at ten inches for the human eye.  This distance is called the resolving power and can be expressed as the relationship of two factors: the wavelength of light and the numerical aperture.

Resolving power  =  ______wavelength_______

numerical aperature

Because this is nothing more than a fraction, and one that needs to be a very small number when divided out, the wavelength must be small and the numerical aperture large.  The reason we commonly use blue-violet filters is because this end of the spectrum is the shortest wavelength that is visible to the human eye.

The light arriving at the specimen is focused from a round base to a point, thus having the volume of a cone (see figure below).  While it is unusual to think of light in volumetric terms; nevertheless, that is the concept of numerical aperture.  It is a mathematical expression of the volume of light within the cone of light delivered to the specimen and objective by the condenser.  The method used to ensure that this cone of light is at maximum volume is called Kohler illumination.

# Setting Up the Microscope for Kohler Illumination

- Begin the alignment at low illumination.

- Use a 10X eyepiece and the 10X objective

- Place an easily visible test slide on the microscope.

- Open the condenser iris diaphragm and partially close the field diaphragm.

- Focus on the test slide.  (With coarse focus control, lower objective as far as it will go, while watching it to see that it does not crush the slide.  DO NOT look through the microscope while lowering objective with coarse focusing knob.  Looking through the microscope, raise objective by means of coarse focus until object comes into view.  Adjust the fine focus.)  Note that the distance between the oculars can be adjusted to your eyes.

f.	Focus the image of the field diaphragm on the test specimen by adjusting the height of the condenser (It will appear polygonal when properly focused).

g.	Center the image of the diaphragm by means of the two set screws located on the condenser.

- Open the field diaphragm until light fills the field.

- When changing to 40X objective, partially close the field diaphragm, then repeat step f-h.  Microscopes should be parfocal, so that only slight adjustments of fine focus should be required when going from one objective to another.

- For oil immersion, first obtain a good image of the desired area with 40X objective.  Swing objective aside, and place a very small drop of immersion oil over bright spot of light on the microscope slide.  Swing 100X objective into place and adjust fine focus.  Always remove oil from the 100X objective with lens paper after use.